## Slide 1: Industrial Instrumentation

## Slide 5: Introduction
Temperature
Viscosity
Density
Flow
Pressure
Humidity
Level
Speed

## Slide 6
Temperature Measurement

## Slide 7
Temperature v/s Heat
Temperature is a direct function of random molecular motion within an object or
a fluid sample.
Heat, by contrast, is an expression of thermal energy transfer

## Slide 8
Molecules in cold state
Molecules in warm state
Molecules in hot state
add external energy
Heat is form of
energy transfer
Remove energy
add external energy

## Slide 9
Illustrating by way of example, a single molecule of gas moving at a constant velocity will have a definite temperature.
Two or three molecules moving at the same speed will have the same temperature, but together represent a greater thermal energy than any one of them considered alone.
Heat is either the reduction or increase of thermal energy by transfer of energy. If these gas molecules happen to collide with slower-moving gas molecules, the faster molecules will lose velocity while the slower molecules will gain velocity.
Thus, the higher-temperature molecules cool down while the lower-temperature molecules warm up: a transfer of heat.

## Slide 10
Temperature
Temperature is measure of the relative hotness or coldness of an object
Temperature is often measured relative to the freezing and boiling points of water

## Slide 12
Temperature is the measure of thermal energy in the body.
Normally measured in degrees [0] using one of the following scales:
1.Fahrenheit.[0F]
2.Celsius or centigrade.[0C]
3.Kelvin. [K]
F = 32 + 9 C / 5
C=K-273.15

## Slide 13: Why measure temperature?
Industry
Healthcare
Meteorology
Science and Technology
Home appliances
Notes: Measuremnt of temperature is important in all phases of human activities including industrial,helath care.home appliances etc..In fact in almost every sectors temperature is the key parameter to be established.

## Slide 14
Different types of temperature measurement methods

## Slide 15
Bi-metal temperature sensors

## Slide 17: Spiral type bimetallic Thermometer
Notes: A classic example of an ambient air temperature thermometer is shown in Figure. The outside edge of the spiral is pinned to the frame and a pointer is connected to the center. As the temperature increases, the spiral winds up deflecting the pointer clockwise. low-expanding Invar (64% Fe, 36% Ni) against high-expanding nickel-iron alloy with chromium or manganese (or with brass)added. Used in ambient temperature measurement devices and air conditioning thermostats.

## Slide 18: Helix Bimetallic Thermometer
Notes: The main parts of a helix bimetallic thermometer are as follows: A bimetallic helix which is fixed at one end to the body of the instrument and free at its other end. To the free end of the bimetallic helix is attached a shaft. One end of the shaft is mounted in friction less arrangement and its other end is connected to a pointer which sweeps over a temperature calibrated circular dial graduated in degrees of temperature.
Used in Industrial process applications

## Slide 19: Liquid in glass thermometers
Filled-bulb systems exploit the principle of fluid expansion to measure temperature. If a fluid is enclosed in a sealed system and then heated, the molecules in that fluid will exert a greater pressure on the walls of the enclosing vessel. By measuring this pressure, and/or by allowing the fluid to expand under constant pressure, we may infer the temperature of the fluid
Glass-stem thermometer

## Slide 20
Mercury or some other liquid (alcohol, pentane) fills the glass bulb and extends into the capillary bore of the stem.
The space above the mercury column to the scaled top is evacuated, but occasionally it may be filled with an inert dry gas, such as nitrogen, to increase the temperature range.
The expansion of pure mercury is 0.01%/°F (0.005%/°C) and very linear; therefore, the volume of the bulb must be about 10,000 times the volume of the capillary between two marks 1°F (0.56°C) apart.
Mercury is most suitable and can be used from its freezing point (−38°F, or −39 °C) up to nearly its boiling point (over 1000 °F, or 538 °C).
Alcohol and a few other hydrocarbons may be used for low temperatures.
Colorfast dyes are usually added to the liquid to increase visibility
Liquid properties
Have large temperature coefficient of expansion.
Temperature- dimensional relationship must be linear
Liquid should accommodate reasonable temperature range
Liquid should be clearly visible when drawn into thread. Alcohol can be used only if a dye is used
Liquid must not stick to the capillary walls.

## Slide 21
Advantages of liquid-in-glass thermometers are their low cost, simplicity, and long life if treated properly.
Disadvantages are difficult reading, confinement to local measurement, and non adaptability to recording or automatic control. They also break very easily.
Mercury filling has been discontinued in most
industrial applications because of health concerns
Click to add text

## Slide 22
FILLED THERMAL SYSTEMS
A filled thermal system is basically a pressure gauge connected by small bore tubing to the bulb acting as the temperature sensor. The whole system is gas-tight, and filled with an appropriate confined gas or liquid under pressure.
There are many different types of filled systems, each having certain peculiarities which give it advantages over others.
The Scientific Apparatus Makers Association(SAMA) has classified filled-system thermometers into four major classes according to filling material

## Slide 23
Liquid-filled capillary-type thermometer

## Slide 24
Class II: vapor system
The pressure element, capillary, and bulb of a Class II system have the filling medium in both the liquid and vapor form.
The interface between the two must occur in the bulb, and this will move slightly with temperature, affecting the pressure.
The pressure within the system is a function of the vapor pressure of the filling fluid at the operating (bulb) temperature.
The filling fluids used include methyl chloride, sulfur dioxide, butane, propane, hexane, methyl ether, ethyl chloride, ethyl ether, ethyl alcohol, and chlorobenzene.
Each has a different vapor pressure–temperature relationship.

## Slide 25
Class II A
This type of system can be used only if the temperature at
the bulb is always above the ambient temperature around the
readout instrument.
As the process temperature rises, more liquid is vaporized,
which increases the vapor pressure in the bulb and in the Bourdon
tube of the readout instrument.

## Slide 26
Class IIB thermal system, all the filling liquid is in the bulb, while the
rest of the system is vapor-filled. This type of configuration
can be used only if the temperature at the bulb is always lower
than the ambient temperature around the capillary
or at the readout device.
If this were not the case, the liquid would start boiling in the bulb
and recondensing in the Bourdon tube.
This transfer of liquid cannot be completed because there is not enough
liquid to fill the Bourdon tube and the capillary;
while the transfer is taking place, the temperature readings are meaningless.
Class II B

## Slide 27
In the Class IID system there are two filling liquids: one is volatile
and is always situated in the bulb, while the other is nonvolatile
and fills the Bourdon tube, capillary, and part of the bulb .
The purpose of the nonvolatile liquid is only to transmit the vapor pressure
at the volatile liquid/vapor interface, which is always confined
within the bulb.
In this design the process temperature can be anything, including the
crossing of the ambient
Class II D

## Slide 28
The operating principle for gas-filled systems is that in a perfect gas confined to a constant volume the pressure is proportional to the absolute temperature

## Slide 29
Compensating techniques
Full compensated
Case compensated

## Slide 30
Invar wire is inserted inside the single active capillary.
Invar is a nickel alloy that has a low coefficient of thermal expansion.
Therefore it is possible to select the diameters of the Invar wire and of the
capillary in such a manner that an increase in ambient temperature will
increase the annular space between the wire and the capillary by just about
the same amount as the volumetric expansion of mercury.
Therefore, a change in ambient temperature around the capillary does not
change the pressure inside the thermal system.
Full compensation with internal wire

## Slide 32
Pneumatic temperature transmitter

## Slide 33
Resistance Temperature Detectors
In 1821 Sir Humphry Davy discovered that the resistivity of metals depends on their temperature.
Resistance thermometry is based upon the increasing electrical resistance of conductors with increasing temperature.
The resistive material may be platinum, nickel, or copper.
Platinum is also available as a deposited film sensor, which is much less expensive, and platinum, nickel, and Balco(70% nickel/30% iron) are all available in foil-type sensors.

## Slide 35
Measurement of RTD resistance

## Slide 36
Two wire RTD

## Slide 37
Two wire RTD

## Slide 38
Three wire RTD

## Slide 39
Three Wire RTD

## Slide 40
Four wire RTD

## Slide 41
RTD connected to transmitter

## Slide 43
Thermistors
NTC thermistors, or NTCs for short, are semiconductors made from specific mixtures of pure oxides of nickel, manganese, copper, cobalt, tin, uranium, zinc, iron, magnesium, titanium, and other metals
PTCs are manufactured from silicon (silistors) or barium,
lead, and strontium titanates with the addition of yttrium,
manganese, tantalum, and silica

## Slide 44
The Steinhart-Hart equation is: 1/T = A + B(lnR) + C(lnR)2 + D(lnR)3 + E(lnR)4… A, B, C, D, and E are the Steinhart-Hart coefficients that vary depending on the type of thermistor used and the range of temperature being detected.

## Slide 45
What is NTC and PTC?
PTC stands for positive temperature coefficient and NTC is negative temperature coefficient. An NTC's resistance will decrease with temperature; where as, the resistance from PTCs will increase. Regardless of which thermistor is being used, it is important that the devices be tested prior using

## Slide 46
The negative exponential function that best describes the
resistance-temperature (R(T)) characteristic of an NTC can
be interpolated using different equations. The Steinhart-Hart
equation
Simplified version

## Slide 47
NTCs can be classified in one of two groups according to the way electrodes are connected to the thermistor body:
Bead-type thermistors and Surface electrode-type thermistors
Bead-type thermistors have platinum lead wires sintered into the sensor body. This group includes bare and ruggedized beads as well as several glass-encased beads, rods, and probes.
Surface-type thermistors have metallized surface contacts and leads. They are cheaper to manufacture than bead-type NTCs. Chip and surface mounts are probably the types more used, but disks, rods, and washers are also in demand.
PTCs can be divided in two groups: thermally sensitive silicon resistors (silistors) and switching PTC thermistors

## Slide 48
Characteristics of PTC

## Slide 49
Applications of PTC thermistors
Self-regulating heaters
If there is a current running through a switching PTC thermistor,
it will auto-stabilize at a certain temperature. It means that if the
temperature is decreased, the resistance will decrease as well,
allowing more current to flow and thus heating the device.
Similarly, if the temperature is increased, the resistance is
increased as well, limiting the current passing through the device,
thus cooling it down.
Over-current protection
Switched PTC thermistors are used as over-current limiters or resettable fuses in various circuits. In the case of an over-current situation, the thermistor body temperature rises and quickly reaches the transition temperature. This results in the resistance of the PTC thermistor sharply rising, limiting current in the circuit
Click to add text

## Slide 50
Measurement of Resistance
Differential measurement
Click to add text

## Slide 51
Two wire thermistor

## Slide 52
Self-heating
In order to measure the resistance of either device, we must pass an electric current through it. Unfortunately, this results in the generation of heat at the resistance according to Joule’s Law.
This dissipated power causes the thermistor or RTD to increase in temperature beyond its
surrounding environment, introducing a positive measurement error.
The effect may be minimized by limiting excitation current to a bare minimum,
but this results in less voltage dropped across the device.

## Slide 53
Thermo-electricity

## Slide 54
Peltier effect
Seebeck effect

## Slide 55
Thermocouple
Thermo for the heat requirement and couple denoting two junctions.
Dissimilar metal junctions
Unknown temperature = (voltage/Seebeck coefficient) + reference temperature

## Slide 59
Thermocouple types

## Slide 60
Dissimilar-metal wires are protected from physical
damage by a stainless steel or ceramic sheath, and
they are often equipped with molded-plastic plugs for quick connection to and disconnection from a thermocouple-based instrument
K type
J type

## Slide 63
Manually interpreting thermocouple voltages

## Slide 64
The (Thermocouple) Thermoelectric Laws
The three fundamental empirical laws behind the accurate measurement of temperature by thermoelectric means are the:
Law of homogeneous materials
Law of intermediate materials
Law of intermediate temperatures.
Notes: The law of intermediate metals states that the introduction of a third metal into the circuit will have no effect upon the EMF generated so long as the junctions of the third metal with the other two are at the same temperature.

## Slide 65
The e.m.f. of a thermocouple depends only on the temperatures of the junctions and is independent of the temperatures of the wires connecting the junctions. This means that the leads connecting the instrument can be exposed to temperature fluctuations without affecting the measurement.
Law of homogeneous materials

## Slide 66
The e.m.f. is not changed when a new conductor is introduced into one of the metals provided that both of the new junctions are at the same temperature.
Law of intermediate materials

## Slide 67
Law 3 states that if a third metal, C, is inserted at one of the junctions, the e.m.f. is unchanged provided that both of the new junctions are at the same temperature. This means that a third material may be used to braze the thermocouple materials together without affecting the measurement

## Slide 68
The law of intermediate metals states that the e.m.f. of materials AC can be predicted if the e.m.f.s of materials AB and BC are known.

## Slide 69
The law of intermediate temperatures states that if two junctions at temperatures T1 and T2 produce a Seebeck voltage V2, and temperatures T2 and T3 produce voltage V1, then temperatures T1 and T3 will produce a voltage V3 = V1 + V2
Law of intermediate temperatures

## Slide 70
Reference junction compensation
Physically fix the temperature of that junction at some constant value so it is always stable.
Fixing the reference junction at the temperature of freezing water is impractical for any
real thermocouple application outside of a laboratory

## Slide 71
The compensating voltage must continuously track the voltage produced by the reference junction. To do this, the compensating voltage source (Vrjc in the above schematic) uses some other temperature-sensing device such as a thermistor or RTD to sense the local temperature at the terminal block where junction J2 is formed and produce a
counter-voltage that is precisely equal and opposite to J2’s voltage (Vrjc = VJ2) at all times

## Slide 73
Software compensation

## Slide 76
Non contact type temperature measurement
Every object in the world radiates thermal energy.
The amount of radiant energy emitted is proportional
to the temperature of an object.
The wavelength at which this radiation occurs depends on the temperature of the material

## Slide 77
Thermal radiation lies in the wavelength region from about 0.1 to 100micrometers

## Slide 78
Radiation pyrometers

## Slide 79
Radiation receiving elements

## Slide 80
Radiation receiving elements
Bolometer
Bolometers give an electric signal as a result of the variation in resistance of a conductor with temperature.
It consist of a thin platinum strip in an evacuated glass vessel, with a transparent window in the infrared range.
Irradiation by the infrared beam produces an increase in resistance of the metal strip, which is measured with a Wheatstone bridge.

## Slide 81: Disappearing Filament Optical Pyrometer

## Slide 83
The shape of the pyrometer is cylindrical and inside parts of the optical pyrometer mainly include eyepiece, power source, absorption screen, and red filter.
An eyepiece and the lens of an object are arranged at both sides of the device.
A battery, millivoltmeter & rheostat are connected to a temperature bulb.
An absorption screen is arranged in the middle of the reference temperature lamp and the objective lens to increase the temperature range and his range can be measured with the help of the device.
The red filter is placed in between the lamp and eyepiece so that the lamp allows simply a narrow band with 0.65mui wavelength.

## Slide 84
Operating Principle
In an optical pyrometer, the temperature measurement principle is used by comparing the brightness. A color disparity with the increase in temperature can be taken as an index of the temperature. This type of pyrometer contrasts the intensity of the generated image through a source of the temperature of the lamp.
The current within the lamp is regulated until the lamp’s brightness is equivalent to the image brightness generated through the source of temperature. When the light intensity of any wavelength depends on the temperature of the radiating object, then the flow of current through the lamp becomes a measure of the temperature source when adjusted.

## Slide 85
Applications
The applications of optical pyrometer include the following.
It is used to measure the temperature of highly heated materials
It is useful to measure furnace temperatures.
It is used in critical process measurements of semiconductor, medical, induction heat treating, crystal growth, furnace control, glass manufacture, medical, etc

## Slide 86
Optical Pyrometer Advantages
The advantages of this pyrometer include the following.
It is used for high temperatures.
It is used to check the distant objects as well as moving the object’s temperature.
Accuracy
It can be measured without connecting with the target.
Less weight
It is flexible and portable.
Optical Pyrometer Disadvantages
Due to the radiation of thermal background, dust, and smoke, the accuracy of this device can be affected.
These do not apply to the temperature measuring of burning gases because they do not emit visible energy.
It is expensive.
Manual type pyrometers are not suitable for evaluating the object’s temperature under 8000C because, at less temperature, the generated energy will be too low.

## Slide 87
An Iron-Constantan thermocouple (TC) is used to measure temperature between 0 o C and 300 oC at which it generates 5268 μV. If the thermal emf is 12500 μV relative to the reference junction at 20 o C. Estimate the measured junction temperature.
Thermocouple:
0 o C to 300 o C 5268 μV
2 0 o C to Th o C 12500 μV
The electro motive force: E=K (Th –Tc )
Where: K is the T / C sensitivity.
5268 =K (300 – 0 )

## Slide 88
Thermowell

## Slide 89
12500 =K (Th –20 )
Th = (12500 x 300-20 x 5268) / 5268 = 691.845 o C

## Slide 90
Thermowell
Thermowell with RTD and temperature transmitter

## Slide 91
The most popular semiconductor temperature sensors are based on the fundamental temperature and current characteristics of the transistor. If two identical transistors are operated at different but constant collector current densities, then the difference in their base-emitter voltages is proportional to the absolute temperature of the transistors. This voltage difference is then converted to a single ended voltage or a current. An offset may be applied to convert the signal from absolute temperature to Celsius or Fahrenheit
Semiconductor temperature sensors
LM-35

## Slide 92
The forward biased voltage across a diode has a temperature coefficient of about 2.3mV/°C and is reasonably linear.
The bias current should be held as constant as possible - using constant current source, or a resistor from a stable voltage source.