## Slide 1: INSTRUMENTATION AND TRANSDUCERS
Click to add text
Click to add text

## Slide 2
Active
They are those which do not require an auxiliary power device to produce an output
Piezo electrical crystal measuring acceleration is an example.

## Slide 3
In this diagram, the crystal is sandwiched between 2 metallic electrodes, a fixed mass is placed on top of the sandwich. The property is that when a force is applied to them, it produces output voltage, The mass is fixed , force is proportional to acceleration, the output voltage is proportional to force and hence to acceleration. This is also called accelerometer, which converts the acceleration to electrical voltage

## Slide 4
Comparison between electrical and mechanical transducer
Electrical transducer | Mechanical transducer
Output signals are electrical in nature | Output signals are mechanical in nature
Electrical quantities are voltage, resistance, capacitance , inductance , phase angle etc. | They are temperature pressure, force, torque, density
Electrical or electronic components are used | Spring elements, bellows are used
Flow measurement isn’t possible | Flow measurement is possible
It is used for pressure and displacement measurement | Orifice, ventury tubes are used for velocity and pressure measurement

## Slide 5
Desirable Characteristics of Transducers
For choosing a transducer for any application the following characteristics should be checked
1) Input characteristics
2) Output characteristics
3) Transfer characteristics
The three characteristics given above are known as the DESIRABLE CHARACTERISTICS OF A TRANSDUCER

## Slide 6
1) INPUT CHARACTERISTICS:
Type of input and operating range. The foremost consideration for the choice of a transducer is the input quantity it is going to measure and its operating range. The transducer should maintain a good resolution throughout its operating range.
b) Loading effects. The transducer, that is selected for a particular application should ideally extract no force, power or energy from the quantity under measurement in order that the latter is measured accurately.

## Slide 7
2) OUTPUT CHARACTERISTICS:
a) Type of electrical output. The types of output which may be available from the transducers may be a voltage, current, impedance or a time function of these amplitudes.
b) Output impedance. The output impedance Zo of a transducer determines to the extent the subsequent stages of instrumentation is loaded.
c) Useful output range. The output range of a transducer is limited at the lower end by noise signals which may hamper the desired input signal. The upper limit is set by the maximum useful input level.

## Slide 8
3) TRANSFER CHARACTERISTICS
a) Transfer function. The transfer function of a transducer defines a relationship between the input quantity and the output quantity.
Transfer function = qo = f(qi)
Where qo is the output quantity and qi is the input quantity
b) Sensitivity.
Static sensitivity
The static sensitivity of an instrument or an instrumentation system is the ratio of the magnitude of the output signal to the quantity to be measured. Its units are counts/volt etc. depending upon the type of input and output.

## Slide 9
FIGURE a)Non feedback(Open loop) system

## Slide 10
The transfer function for the desired inputs in a non-feedback (open loop) system is GD and the desired input is rD. The output due to desired input C=GDrD.

## Slide 14
Suppose, due to the modifying input, like changes in ambient temperature, the transfer function changes by ∆GD and therefore modifying transfer function is GMD=GD+ ∆GD.

## Slide 15
Null-type and deflection-type instruments
In deflection type instruments, the value of the quantity being measured is displayed in terms of the amount movement of a pointer. The pressure-measuring device shown is an example of a deflection type instrument.
An alternative type of pressure gauge is the deadweight gauge which is a null-type instrument. Here, weights are put on top of the piston until the downward force balances the fluid pressure. Weights are added until the piston reaches a datum level, known as the null point. Pressure measurement is made in terms of the value of the weights needed to reach this null position.

## Slide 16
Sensitivity
Sensitivity is the ability of the transducer to generate
an output for a given change in input.
Sensitivity =Change in output / Change in input
E.g. A thermocouple that increases output
voltage by 3mV per degree Celsius temperature
change has a sensitivity of 3mV/ C

## Slide 17
Analogue and digital instruments
An analogue instrument gives an output that varies continuously as the quantity being measured changes. The output can have an infinite number of values within the range that the instrument is designed to measure. The deflection-type of pressure gauge is an example of an analogue instrument.
As the input value changes, the pointer moves with a smooth continuous motion. Whilst the pointer can therefore be in an infinite number of positions within its range of movement, the number of different positions that the eye can discriminate between is strictly limited, this discrimination being dependent upon how large the scale is and how finely it is divided.

## Slide 18
Analogue and digital instruments
A digital instrument has an output that varies in discrete steps and so can only have a finite number of values. The rev counter is an example of a digital instrument. A cam is attached to the revolving body whose motion is being measured, and on each revolution the cam opens and closes a switch. The switching operations are counted by an electronic counter. This system can only count whole revolutions and cannot discriminate any motion that is less than a full revolution

## Slide 19
Sensitivity of a sensor is defined as the
change in output for a given change in input, usually a unit change in input.
Sensitivity represents the slope of the transfer function.
Also is used to indicate sensitivity to other environment that is not measured.

## Slide 20
Sensitivity = dqo/dqi
When calibration curve is linear, the sensitivity of the instrument can be defined as the slope of the calibration curve. In this case the slope will be constant over the entire range of the instrument. However if the curve is non linear, the slope varies with input shown in the fig.

## Slide 21
c) Hysteresis. It is a phenomenon which depict different output effect while loading and unloading whether it’s a mechanical system or an electrical system or any system. It’s a non coincidence of loading and unloading curves

## Slide 22
Hysteresis
• Hysteresis is the deviation of the sensor’s output at any given point when approached from two
different directions

## Slide 23
• When the input of the instrument is slowly varied from 0 to full scale and the back to 0, its output varies as shown in fig.(a)

## Slide 24
• In the case of the instruments which are used on both sides of 0 (+ve & -ve applied input) the variation of output is shown in fig.(b)

## Slide 25
• At constant Columb’s friction, the input-output relationship are like the ones shown in fig.(c) & fig.(d)

## Slide 26
Statement. The law states that for two dry solid surfaces sliding against one another, the magnitude of the kinetic friction exerted through the surface is independent of the magnitude of the velocity (i.e., the speed) of the slipping of the surfaces against each other.

## Slide 27
• The maximum input hysteresis & maximum output hysteresis is shown in fig.(e)

## Slide 28
Accuracy
Accuracy can be expressed as a comparison of the static error of the transducer compared to the actual
value (at full scale) expressed as a percentage of full
scale. (Accuracy may also be expressed in other
ways.)
% Accuracy
(Measured value – Actual value) x 100/ Actual value
=
E.g. A temperature transducer that reads 102 C at full scale, when the temperature is 100 C, has an accuracy equal to 2% of full scale.

## Slide 29
Range
The highest and lowest values that the
transducer is designed to measure.
E.g. A Temperature transducer may have a range of –50 C to +50 C
Span
The difference between the upper and lower values the transducer is designed to measure.
•E.g. A Temperature transducer that has a range
of –50 C to +50 C has a span of 100 C

## Slide 30
Range and Span
• Range: lowest and highest values of the
stimulus
• Span: the arithmetic difference between the
highest and lowest values of the input that
being sensed.
• Input full scale (IFS) = span
• Output full scale (OFS): difference between
the upper and lower ranges of the output of
the sensor.
• Dynamic range: ratio between the upper and
lower limits and is usually expressed in db

## Slide 31
Range and Span (Example)
• Example: a sensors is designed for
: -30 °C to +80 °C to output 2.5V to 1.2V
• Range: -30°C and +80 °C
• Span: 80- (-30)=110 °C
• Input full scale = 110 °C
• Output full scale = 2.5V-1.2V=1.3V
• Dynamic range=20log(140/30)=13.38db

## Slide 32
Errors and Accuracy
•Errors: is the difference between the result ofthemeasurement and the true value of the quantity being measured
error= measured value –true value
• As a percentage of full scale (span for
example) error is calculated as;
e = Dt/(tmax-tmin)*100
where tmax and tmin are the maximum and minimum values the device

## Slide 33
Errors and Accuracy Example:
• Accuracy: is the extent to which the measured value might be wrong and normally expressed in percentage
• Example: A thermistor is used to measure
temperature between –30 and +80 °C and produce an output voltage between 2.8V and 1.5V. Because of errors, the accuracy in sensing is ±0.5°C. so the measured value may be high than or lower than by 0.5 °C

## Slide 34
STATIC ERROR
The most important characteristic of an instrument is its accuracy which is the agreement of the instrument reading with the true value of the quantity being measured. The accuracy of an instrument is measured in terms of the error. Static error is defined as the difference between the measured value and true value of the quantity.
δA = Am – At
Ɛ0 = δA
Ɛr = (δA / At) = (Ɛ0 / At)
At = Am /(1 + Ɛr)
Since, Ɛr << 1
At = Am (1 - Ɛr)

## Slide 35
DYNAMIC CHARACTERISTICS
Dynamic characteristics of a measuring system relates to its performance when the measurand is a function of time.
The dynamic response of a measurement system when subjected to dynamic inputs which are function of time depends very much on its own parameters apart from the nature and complexity of the function.
Thus the dynamic response of a measurement system consists of two components, one due to its characteristic parameters and the other due to the nature of the input function.

## Slide 37
DYNAMIC ERROR
It is defined as the algebraic difference between the indicated / recorded value and its true value at any instant when the measuranda is a measure of time.
It is a function of time and this error is zero only for the zero order system.
For measuring the higher order, their output signals consist of two components, one pertaining to the transient state, other to the steady state.

## Slide 38
Linearity
Linearity refers to the change in output compared to the change in input. If the change in output is proportional to the change in input, the transducer is said to be linear.

## Slide 39
tNon linearity means that the output is not constant with respect to the input signal.
o the input signal.
Out
Input
Input
Input
Output

## Slide 41
Null Type Instrument.
Definition: An instrument in which zero or null indication determines the magnitude of measured quantity such type of instrument is called a null type instrument.
It uses a null detector which indicating the null condition when the measured quantity and the opposite quantity are same.

## Slide 42
In deflection type instruments, the value of the quantity being measured is displayed in terms of the amount movement of a pointer. The pressure-measuring device shown is an example of a deflection type instrument. An alternative type of pressure gauge is the deadweight gauge which is a null-type instrument.

## Slide 43
Null-type and deflection-type instruments
In deflection type instruments, the value of the quantity being measured is displayed in terms of the amount movement of a pointer. The pressure-measuring device shown is an example of a deflection type instrument.
An alternative type of pressure gauge is the deadweight gauge which is a null-type instrument. Here, weights are put on top of the piston until the downward force balances the fluid pressure. Weights are added until the piston reaches a datum level, known as the null point. Pressure measurement is made in terms of the value of the weights needed to reach this null position.

## Slide 45
Analog and Digital Instruments:
An analogue instrument gives an output that varies continuously as the quantity being measured changes. The output can have an infinite number of values within the range that the instrument is designed to measure. The deflection-type of pressure gauge is an example of an analogue instrument.
An instrument whose output is in digital form is when there is a need to be interfaced to a control computer. Analogue instruments must be interfaced by an analogue-to-digital (A/D) converter.

## Slide 47
A cam is a rotating or sliding piece in a mechanical linkage used especially in transforming rotary motion into linear motion

## Slide 48
SELECTION CRITERIA OF THE TRANSDUCERS
  Operating principle
Sensitivity
Operating range
Accuracy
Errors
Environmental capability
Insensitive to unwanted Signal
 Stability
Cost

## Slide 49
Generalized Measurement System

## Slide 50
An instrument may be defined as a device or a system which is designed to maintain a functional relationship between the prescribed properties of physical variables and must include ways and means of communication to a human observer.

## Slide 51
The performance of a measurement system can be described in terms of static and dynamic characteristics. Most of the measurement systems contain three main functional elements: i) Primary sensing elements ii) Variable conversion elements iii) Data presentation elements

## Slide 52
1.Primary sensing element: The quantity under measurement makes its first contact with the primary sensing element of a measurement system.
The measurand is first detected by a primary sensor. This is done by a transducer. 2.Variable conversion element: The output of a primary sensory element may be an electrical signal of any form. It may be voltage, frequency or other electrical parameters. Sometimes this output is not suited to the system for the instrument to perform the desired function. It may be necessary to convert this output to some other suitable form. Eg. Suppose output is analog form, then we have to convert it into digital form.

## Slide 53
3.Variable manipulation element:
The function is to manipulate the signal presented to it preserving the original nature of the signal. Manipulation means change in numerical value of signal. 4.Data transmission element: When the functional elements of the measuring system are spatially separated then it becomes necessary to transmit signals from one element to another. This function is performed by data transmission element. It is an essential functional element where remote control operation is desired.

## Slide 54
Comparison between electrical and mechanical transducer
Electrical transducer | Mechanical transducer
Output signals are electrical in nature | Output signals are mechanical in nature
Electrical quantities are voltage, resistance, capacitance , inductance , phase angle etc. | They are temperature pressure, force, torque, density
Electrical or electronic components are used | Spring elements, bellows are used
Flow measurement isn’t possible | Flow measurement is possible
It is used for pressure and displacement measurement | Orifice, ventury tubes are used for velocity and pressure measurement